INSTITUTE OFMICROBIAL TECHNOLOGY (IMTECH),

(Council of Scientific & Industrial Research)

Sector 39-A, Chandigarh -160036 (India)

## Ph: 0091-172-2690056/2695225/6/7 ext. 331/334, fax:2 690585/2690632, 2636683

## e-mail: purchase@imtech.res.in

| No: PUR/2010-11/IND14501/KK | Date: 27/12/2010 |
| --- | --- |

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | Automated controlled atmosphere isolater/glove box, 1 |  | 1 |

Important - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 11/01/2011 opened on 12/01/2011 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation.

Yours faithfully,

Stores & Purchase Off

Specifications for inviting quotation for purchase of an anaerobic glove box

Minimum Specifications: Glove Box must have the following specification

1. Chamber'shall have automatic and manual vacuuming and gassing facility. The chamber shall have dimensions of approximately 800mm x 1000mm x 700mm (H x W x D) or better with four sleeves to facilitate work of two person simultaneously. The chamber with an attached antechamber for large equipment entry port (with automatic and manual vacuuming and gassing facility) will be preferred. The box shall have special protection layer for chemical & scratch resistance. The box shall be equipped with fluorescent lamp and shelves inside. This chamber shall maintain an oxygen and moisture controlled environment.

2. This instrument shall be equipped with a dust filter to prevent contamination. The preferred filter will be with one gas inlet/one gas outlet filter. All the glove ports shall have O-ring sealing. Thick gloves resistant to air seepage are essential.

3. Chamber should be equipped to maintain controlled pressure inside the box adjustable from -4mbar to +4mbar. But glove boxes with higher adjustable pressures will be preferred.

4. Automatic box pressure should be controlled with foot pedal operation that shall be water proof. A Cylindrical antechamber with automatic and/or manual control of pressure/vacuum will be preferred. The preferred antechamber will have pressure gauge, manometer analogue display vacuum/refill process-manual operation via hand valves or automatic or preferably both.

5. Chamber shall have closed loop gas re-circulation preferably via single filter purifier bed. High re-generable oxygen removal capacity and moisture removal capacity will be preferred. Circulation unit should consist of integrated blower vacuum tight and a high flow rate. Vacuum pump will preferably be type rotary vane pump.

6. Electro-pneumatic, valves of high vacuum will be preferred for the main box and high vacuum valves will be preferred for the Antechamber. 7. Manual regeneration of purifier material is essential and automatic regeneration of purifier material through regeneration program will be preferred.

8. Oxygen and moisture probe should be provided with the machine to analyze the levels of oxygen and moisture (with digital display) inside the chamber. Glove box with the integrated sophisticated sensor for low level of oxygen and moisture will be preferred.

9. Glove box should be equipped    with     water  and/or solvent trap ,for removing any hazardous organic solvents.

- Digital display of the inside chamber temperature on control panel board will be preferred.

- Provision must be available to use more than one gas at a time.

TERMS & CONDITIONS

Commercial

1) Cost                  : 	Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment.

2) Insurance         :	Consignment should be insured through National Insurance               Company of India or is associate in your country.

3) Freight              :	The dispatch of the material should be through Air India preferably.

4) Agency             : 	Mention the   percentage   of  Indian  Agency   Commission  payable  to

Commission                    commission Indian Agent in INR and also mention the name of your   authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end.

5) Validity             : 	The validity of your pro forma invoice should be less than 180 days from the date of opening.

6) Warranty           : 	Mention the warranty period of equipment after satisfactory   installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above

General Condition

- Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market.

- Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization.

- Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years.

- Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required.

- Provide the list of the users in India of the same equipment along with their contact / email address.

- Enclose a copy of agreement with your authorized Indian Agent.

- 10% Performance Bank Guarantee must be submitted if order value beyond 25.00 lakhs.

FAX : 0091-172-2690056, 2690632			                                   email : purchase@imtech.res.in

TEL  : 0091-0172-2690056, 2636680 - 94     	                                            website : www.imtech.res.in