| No: PUR/2012-13/IND22918/null | Date: 18/09/2012 |
| --- | --- |

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | Gel Dryer, 1 |  | 1 |

Important - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 1/10/2012 opened on 3/10/2012 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation.

Yours faithfully,

Stores & Purchase Officer.

Detailed specification for gel drying system

1. Gel Dryer should Includes Porous Gel Support and transparent sealing gasket.

2. It should have lid mounted heating plate to heat the gel from top

3. Filter Paper, cellophane membrane backing

4. Drying Surface should be 35x45cm or more.

5. System should have accommodation capacity of 16mini(8X7cm) gels, 9 midi (13.3X8.7cm) gels,2 large(20x20cm) gels , 1sequencing gel for uniform drying.

6. There should be a polymeric porous gel support made of polypropylene or strong polymer and should have 120 um average pore size.

7. Maximum heating temperature should be 90 degC or more.

8. Vacuum assembly should be applied at the bottom to collect the vaporsm and minimum pump pressure of 25” should be recommended to hold the gel firmly.

9. There should be separate at least three programmable cycles for PAGE, sequencing and gradient gels.

10. Operating time 5 min to 9 hr 55 Min; it should have operating time setting with at least 1-2 hr timer.

11. The instrument should have LED display for temperature and time.

12. System should be certified by CE, EN 61010-1 or ISO.

13. The company must prvide 2 years comprehensive warranty.