| No: PUR/2012-13/IND23680/null | Date: 19/11/2012 |
| --- | --- |

To

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | Water Purification System, Attached |  | 2 |

Important - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 4/12/2012 opened on 5/12/2012 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation. - EMD Rs ...... must be deposited along with quotation.

4/12/20125/12/2012

TERMS & CONDITIONS Commercial - Cost : Mention the price of the material in the FOB/FCA/CIF including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment. - Cost of Packing, Forwarding, Freight, Transportation, Insurance charge may please be mentioned invariably, failing which offer is liable to be rejected. - Agency Commission : Mention the percentage of Indian Agency Commission payable to commission Indian Agent in INR and also mention the name of your authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end. - Validity : The validity of your pro forma invoice should be less than 180 days from the date of opening. - Warranty : Mention the warranty period of equipment after satisfactory installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above - Taxes For Indigenous products, please quote Sales Tax/VAT/Service Tax/Freight/Other levies etc., if any, clearly in your quotation. General Condition - Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market. - Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization. - Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years. - Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required. - Provide the list of the users in India of the same equipment along with their contact / email address. - Enclose a copy of agreement with your authorized Indian Agent. - FAX : 0091-172-2636683, 2690632 email : purchase@imtech.res.in - TEL : 0091-0172-6665363 website : www.imtech.res.in

Yours faithfully,

Stores & Purchase Officer

Stage1: Pre Filter

5 Micron and 1 Micron depth filters with Low Voltage DC Pump.

Stage2: Analytical Grade Water System(Type II)

Output should meet or exceed Type II with the following Technical Specifications:

Resitivity  5 Megh Ohms

TOC Levels less than 30 ppb

Flow Rate 10 Ltrs/Hr

The System should have ideally a three stage Purification system:

Stage 1: Pretreatment Cartridge with anti scaling compounds, activated carbon filter and 1 Micron Particulate filter, and compatible with Feed Water Quality of SDI levels up to 12 and total Chlorine level of 3 ppm and conductivity of 2000 micro Siemens/cm.

Stage 2: RO Membrane with conductivity cells before and after the Membrane to 95-99 % rejection of Inorganic Ions, 99% rejection of all Dissovled organic substances.

Stage 3: Electro De-ionization module (EDI Module), with  Carbon Beads at cathode to avoid scaling.

The system should have the following:

In built display to ensure the system parameters

Auto diagnostic system

Automatic Cleaning, Rinsing, and Flush mode.

Indication for poor quality water delivery.

Stage 3: Storage Reservoir

100 Litre PE tank with inbuilt sensor to prevent overflow

Stage-4: Ultrapure Water Purification System

Should have mixed bed ion exchange and activated carbon cartridge and  a 0.22 micron filter.

Recirculation of water

Dispensing height should be adjustable to any glassware.

Volumetric dispensing from 0.5 ml to 10L.

Display for resistivity, level of water in reservior, volume dispensed and other alarms.

UltraPure (Type I) water quality

Ultrapure Water (Type 1) Flow Rate (L/min)		1.2

Ultrapure Water Resistivity (MΩ·cm at 25°C)	18.2

Microorganisms (cfu/mL)				<0.1

Particulates < 0.22 μm ( / mL)			< 1

TOC (ppb) 					< 10

General conditions:

The supplier must supply following along with quotation:

Compliance statement in tabular form giving details as per specifications given in the tender document

Literature in support of compliance statement

Details of AMC charges after warranty period

List and price of minimum recommended spares for 5 years.

List of users