Top of Form

| No: PUR/2013-14/IND28872/KK | Date: 26/09/2013 |
| --- | --- |

To

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | HPLC with UV & RI detectors, 1 |  | 1 |

Important - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 08/10/2013 opened on 09/10/2013 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation. - EMD Rs ...... must be deposited along with quotation.

08/10/201309/10/2013

TERMS & CONDITIONS Commercial - Cost : Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment. - Cost of Packing, Forwarding, Freight, Transportation, Insurance charge may please be mentioned invariably, failing which offer is liable to be rejected. - Agency Commission : Mention the percentage of Indian Agency Commission payable to commission Indian Agent in INR and also mention the name of your authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end. - Validity : The validity of your pro forma invoice should be less than 180 days from the date of opening. - Warranty : Mention the warranty period of equipment after satisfactory installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above - Taxes For Indigenous products, please quote Sales Tax/VAT/Service Tax/Freight/Other levies etc., if any, clearly in your quotation. General Condition - Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market. - Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization. - Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years. - Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required. - Provide the list of the users in India of the same equipment along with their contact / email address. - Enclose a copy of agreement with your authorized Indian Agent. - FAX : 0091-172-2636683, 2690632 email : purchase@imtech.res.in - TEL : 0091-0172-6665363 website : www.imtech.res.in

Yours faithfully,

Stores & Purchase Officer.

Bottom of Form

TECHNICAL SPECIFICATIONS ANALYTICAL HPLC

1.            Solvent Delivery System for Analytical flow rates

- It should be High Pressure Binary Gradient two Pump System

- The flow rate should be 0.001 to 10 ml/min

- Flow rate precision should be less than 0.1% RSD

- Flow rate accuracy should be 1%

- Pressure setting range should be 1-40 MPa

- Maintenance kit, reservoir tray (with 2 solvent bottles complete with fittings)  & automatic rinsing kit must be supplied

- It must be capable of upgrading to low pressure Quaternary gradient operation if required.

- It must have a leak sensor as safety feature

- Gradient Mixer should be provided.

- 2.	UV-VIS Spectrophotometer Detector

- Light Source: Deuterium (D2)Lamp/Tungsten Lamp/ Halogen Lamp

- Wavelength Range: 190nm to 700 nm

- Band width, Slit width: 8 nm

- Wavelength precision: ± 0.1nm

- Wavelength accuracy: ± 1.0 nm

- Noise: +0.25X10 -5 AU (0.5x10 -5 AU)

- Drift : 1x10-4 AU/hour

- Cell: Optical wavelength: 10mm, Capacity: 12 UL , withstand pressure: 12MPa

- 3.	Sample Injectors :

Rheodyne Manual with 20 ul loop, Mounting Plate and syringes 25 ul, 250 ul.

4.           Refractive Index Detector

- Refractive Index range : 1-1.75 RIU

- Noise Level should be less than 2.5 x 10-9 RIU

- Drift should be less than 1 x 10-7 RIU/hour

- Range : A mode : 0.01-500 x 10-6 RIU

P and L modes: 1 - 5000 x 10-6 RIU

- Response : 0.05-10 seconds, 10 step selection

- Zero point adjustment : Auto, Auto optical, and Fine

- Flow Cell Volume : 9 uL

- Ambient Temp. : 4-35 Deg. C

- .  	Computer and Software

- All in one Desktop system with i 7 processor, 8GB RAM, 00 GB hard Disk, Microsoft office, windows 7 compatible operating system.

- It should cover full one-point digital instrument control, qualitative and quantitative processing, report creation and self-diagnosis

- The software should allow automatic execution of system checks, auto-purge and baseline checks

- Sample schedule wizard function should be standard

- There should be an on-line help function context sensitive

- System suitability, System security as well as System check functions must be provided which comply with Good Laboratory Practice (GLP) and Regulatory Conformity

- 6.	Installation and Operational Qualification (IQ/OQ) Quoted separately

- 7. 	Warranty:  minimum two years of standard warranty should be quoted