Top of Form

| No: PUR/2013-14/IND28881/KK | Date: 04/10/2013 |
| --- | --- |

To

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | HPHV Double Door Steam Sterilizer (Under Buyback) , Pls see attachment |  | 1 |

Important - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 22/10/2013 opened on 23/10/2013 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation. - EMD Rs ...... must be deposited along with quotation.

22/10/201323/10/2013

TERMS & CONDITIONS Commercial - Cost : Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment. - Cost of Packing, Forwarding, Freight, Transportation, Insurance charge may please be mentioned invariably, failing which offer is liable to be rejected. - Agency Commission : Mention the percentage of Indian Agency Commission payable to commission Indian Agent in INR and also mention the name of your authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end. - Validity : The validity of your pro forma invoice should be less than 180 days from the date of opening. - Warranty : Mention the warranty period of equipment after satisfactory installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above - Taxes For Indigenous products, please quote Sales Tax/VAT/Service Tax/Freight/Other levies etc., if any, clearly in your quotation. General Condition - Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market. - Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization. - Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years. - Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required. - Provide the list of the users in India of the same equipment along with their contact / email address. - Enclose a copy of agreement with your authorized Indian Agent. - FAX : 0091-172-2636683, 2690632 email : purchase@imtech.res.in - TEL : 0091-0172-6665363 website : www.imtech.res.in

Yours faithfully,

Stores & Purchase Officer.

- TECHNICAL SPECIFICATION FOR HPHV STEAM STERILIZER

- Sterilizer Type: Horizontal Rectangular Double Door type.

- Door Type: Hinge Door / Horizontal Sliding Door (With Door and Process Interlock) fitted with silicone rubber door gasket. Bidders should quote separately for both the hinge type door as well as for horizontal sliding door.

- Capacity / Size: 1200 (W) x 1200 (H) x 2100 (D) mm

- Construction:  	i.	Chamber: SS 316

ii.	Door: SS 304

iii.	Jacket: SS 304

iv.	Insulation Cover: SS 304

v.	Piping: SS 304

- Chamber / Door  Insulation: Glass Wool

- Mounting: On stainless steel legs.

- SS 304 sheet covering for surrounding wall on the sterile side should be provided

- Vacuum pump: Water ring vacuum pump

- Steam condenser: Shell and Tube type condenser of SS 304

- Vacuum Breaker: 0.2-micron vacuum break filter

- Steam Generator: Fabricated from SS 316 sheets with industrial immersion heaters, automatic pressure control and other safety features like low water cut off, safety valve gauge glass etc.

- Validation Port for insertion of 16-20 flexible temperature sensors.

- Process Control: Programmable Logic Controller with user friendly HMI,  Printer Port with Printer

- Temperature and Pressure Sensor: Temperature sensor should have least count of 0.10 C.

and Pressure sensor should have least count of 0.002 bar.

- Manual Back Up: The control system shall be provided with manual back up in case of PLC failure.

- Safeties And Alarms:

i. 	Safety Valve:  safety valve shall be provided in the jacket

ii.	Baffle: For effective distribution of steam in the chamber

iii. Plug Screen: The chamber end of the drain line shall have a removable plug screen

iv. Audiovisual alarm for low/high temperature and pressure, door not properly closed, vacuum pump trip etc.

- Must have a Local service engineer for repairing within 24 Hours of complaint.

- National or International Certification/validation is mandatory

- Complete installation and demonstration of performance of the equipment

- Onsite training of personnel to operate the system

- Warranty period should be of 2 Years.

- The bidder should have executed at least one similar order successfully (in India or abroad) during the preceding three financial years. The details should be incorporated in the performance statement form along with documentary evidence.

- Other things being equal, preference shall be given to firms who or his principal has supplied and installed similar system at any CSIR/ ICAR/ ICMR/DAE/ DRDO/ DST/DBT/other Govt. or autonomous research Labs in India.

- Buy back of our existing autoclave

IMTECH has a bulk sterilizer which was installed in 1992 of following specifications:

Model No. 42- LB/DS/DD, double door

Make- Nat steel Equipment PVT. LTD.

Operating Pressure 	1.2 KG/CM2	Temperature: 1210C

Capacity / Size: 1200 (W) x 1070 (H) x 2100 (D) mm.

- Institute invites the bidders to quote a separate buy back price for the old autoclave.

- The Purchaser reserves the right to place the order with or without ‘buy back’ option.

- If required the condition of old buy back autoclave may be examined by the bidder before submission of its bid.

Quotation/Tenders are invited in Two Bid System i.e. 1) Technical Bid  2) Price Bid

Specific instruction for submission of Price Bid

1.  Price for New System as per our Tender Specification which includes cost of the system as per specification, specification giving details of taxes duties etc.

2. Price Offered for existing Autoclave (detail as per tender) – Bidder may examine the old autoclave before submission of the price bid.

1Bottom of Form