INSTITUTE OF MICROBIAL TECHNOLOGY (IMTECH),

(Council of Scientific & Industrial Research)

Sector 39-A, Chandigarh -160036 (India)

## Ph: 0091-172-2690056/2695225/6/7 ext. 331/334, fax:2 690585/2690632, 2636683

## e-mail: purchase@imtech.res.in

| No: PUR/2010-11/IND16086/KK | Date: 17/02/2011 |
| --- | --- |

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | Anaerobic chmaber to work on anaerobic microorganisms, 1 |  | 1 |

Important Important : Quote in Two bid i.e. (1) Tech. Bid (2) Price Bid - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 01/03/2011 opened on 03/03/2011 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation.

Yours faithfully,

Stores & Purchase O

Specifications for purchase of an anaerobic glove box

Minimum Specifications: Glove Box must have the following specification

- Chamber shall have automatic and manual vacuuming and gassing facility. The chamber (inside) shall have dimensions of approximately 650 mm x 1000 mm x 700 mm (H x W x D) or better with two sleeves to facilitate work of two person simultaneously. The chamber with an attached antechamber for large equipment entry port (with manual vacuuming and gassing facility) will be preferred. The box shall have special protection layer for chemical & scratch resistance.

- The box shall be equipped with fluorescent lamp and shelves inside. This chamber shall maintain anaerobic and moisture controlled environment. Electric strip with multiple socket should be provided inside chamber.

- This instrument shall be equipped with a dust filter to prevent contamination. The preferred filter will be with one gas inlet/one gas outlet filter. All the glove ports shall have O-ring sealing. Thick gloves resistant to air seepage are essential and intact gloves shall enable to work inside chamber.

- Chamber should be equipped to maintain controlled pressure inside the box adjustable from –4mbar to +4mbar. But glove boxes with higher adjustable pressures will be preferred.

- A heating element should be available inside the chamber to sterilize the caps of vials and controlled with foot pedal operation that shall be water proof. A Cylindrical antechamber with manual control of pressure/vacuum will be preferred. The preferred antechamber will have pressure gauge, manometer analogue display vacuum/refill process-manual operation via hand valves or automatic or preferably both.

- Chamber shall have closed loop gas re-circulation preferably via single filter purifier bed.  High re-generable oxygen removal capacity and moisture removal capacity will be preferred. Circulation unit should consist of integrated blower vacuum tight and a high flow rate.  Vacuum pump will preferably be type rotary vane pump.

- Electro-pneumatic, valves of high vacuum will be preferred for the main box and high vacuum valves will be preferred for the Antechamber.

- Manual regeneration of purifier material is essential and automatic regeneration of purifier material through regeneration program will be preferred.

- Oxygen and moisture probe should be provided with the machine to analyze low levels of oxygen and moisture (with digital display) inside the chamber.

- Digital display of the inside chamber temperature on control panel board will be preferred.

- Provision must be available to use more than one gas at a time.

- A minimum of 2 years guarantee must be provided.

TERMS & CONDITIONS

Commercial

1) Cost                  : 	Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment.

2) Insurance         :	Consignment should be insured through National Insurance               Company of India or is associate in your country.

3) Freight              :	The dispatch of the material should be through Air India preferably.

4) Agency             : 	Mention the   percentage   of  Indian  Agency   Commission  payable  to

Commission                    commission Indian Agent in INR and also mention the name of your   authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end.

5) Validity             : 	The validity of your pro forma invoice should be less than 180 days from the date of opening.

6) Warranty           : 	Mention the warranty period of equipment after satisfactory   installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above

General Condition

- Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market.

- Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization.

- Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years.

- Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required.

- Provide the list of the users in India of the same equipment along with their contact / email address.

- Enclose a copy of agreement with your authorized Indian Agent.

- 10% Performance Bank Guarantee must be submitted if order value beyond 25.00 lakhs.

FAX : 0091-172-2690056, 2690632			                                   email : purchase@imtech.res.in

TEL  : 0091-0172-2690056, 2636680 - 94     	                                            website : www.imtech.res.in

INSTITUTE OF MICROBIAL TECHNOLOGY (IMTECH),

(Council of Scientific & Industrial Research)

Sector 39-A, Chandigarh -160036 (India)

## Ph: 0091-172-2690056/2695225/6/7 ext. 331/334, fax:2 690585/2690632, 2636683

## e-mail: purchase@imtech.res.in

| No: PUR/2010-11/IND16142/KK | Date: 17/02/2011 |
| --- | --- |

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | refigerated incubator shaker, - | - | 2 |

Important Important : Quote in Two bid i.e. (1) Tech. Bid (2) Price Bid - Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. - Please quote strictly as per our specifications. - Please attach compliance sheet - Proforma Invoice received after due date will not be considered under any circumstances. - Please note that your Quotation should reach us latest by 01/03/2011 opened on 03/03/2011 - Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation.

Yours faithfully,

Stores & Purchase Officer.

Specifications:

# Shaking Range: 25 to 400rpm with +/- 1 rpm

# Orbit diameter at least 2.5cm

- Temperature :  15°C below ambient to 80°C within +/-0.1oC  and uniformity +/-0.30oC

- Adjustable height shelf for incubations

- Interface to computer

- Audible and Visible alarm should be there

- Universal Platform should have capability to hold at least 18 flasks of 250ml  or more

8) Universal platform can hold the flasks from 50ml to 6L

Clamps to be supplied:  250ml (10nos), 500ml (10nos), 1000ml (5 numbers), 2000ml (5 numbers)

9) Shaker should be supplied with Inbuilt Programmable Ultra Violet Light

10) Shaker should be stackable one over other

11) Comprehensive warranty for two years

Optional:

Shaker should have upgradable provision for Gassing manifold, Photosynthesis etc.

TERMS & CONDITIONS

Commercial

1) Cost                  : 	Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment.

2) Insurance         :	Consignment should be insured through National Insurance               Company of India or is associate in your country.

3) Freight              :	The dispatch of the material should be through Air India preferably.

4) Agency             : 	Mention the   percentage   of  Indian  Agency   Commission  payable  to

Commission                    commission Indian Agent in INR and also mention the name of your   authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end.

5) Validity             : 	The validity of your pro forma invoice should be less than 180 days from the date of opening.

6) Warranty           : 	Mention the warranty period of equipment after satisfactory   installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above

General Condition

- Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market.

- Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization.

- Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years.

- Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required.

- Provide the list of the users in India of the same equipment along with their contact / email address.

- Enclose a copy of agreement with your authorized Indian Agent.

- 10% Performance Bank Guarantee must be submitted if order value beyond 25.00 lakhs.

FAX : 0091-172-2690056, 2690632			                                   email : purchase@imtech.res.in

TEL  : 0091-0172-2690056, 2636680 - 94     	                                            website : www.imtech.res.in